## Sheet1 (hidden)
|  | Full/Part Time | Clinical | Part Time | Indirect | Inflation
|  | Benefit Eligible | Faculty | Non-Benefit Eligible | Cost Rate | Rate
Current | 2008 | 30.2% | 15.8% | 9.7% | 52% | 0% |  |  |  | Month
Current | 2009 | 30.2% | 15.8% | 9.7% | 52% | 0% |  | Start Date | #REF! | #REF!
Current | 2010 | 32.2% | 16.8% | 10.5% | 53% | 0% |  | End Date | #REF!
Current | 2011 | 32.2% | 17.2% | 10.6% | 54% | 0% |  | No. of Months | #REF! |  |  | #REF!
Current | 2012 | 32.2% | 16.8% | 10.5% | 54.5% | 0%
Current | 2013 | 32.5% | 16.8% | 10.2% | 54.5% | 3%
[8 empty] | Year 1 |  | Year 2 |  | Year 3 |  | Year 4 |  | Year 5
|  | Month | Current FY | Next FY |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 1 | 6 | 6 |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 2 | 5 | 7 |  |  | IDC | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 3 | 4 | 8 |  |  | Full Time Benefit Eligible | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 4 | 3 | 9 |  |  | Clinical Faculty | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 5 | 2 | 10 |  |  | Part Time Non-Benefit Eligibile | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 6 | 1 | 11 |  |  | Inflation Rate | 0% | #REF!
|  | 7 | 12 | 0
|  | 8 | 11 | 1 |  |  |  | Months Remaining |  | Months Remaining |  | Months Remaining |  | Months Remaining |  | Months Remaining
|  | 9 | 10 | 2 |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 10 | 9 | 3 |  |  |  | Weighted
|  | 11 | 8 | 4 |  |  | IDC | #REF! |  | #REF! |  | #REF! |  | #REF! |  | #REF!
|  | 12 | 7 | 5 |  |  | Full Time Benefit Eligible | 33% |  | #REF! |  | #REF! |  | #REF! |  | #REF!
[7 empty] | Clinical Faculty | #REF! |  | #REF! |  | #REF! |  | #REF! |  | #REF!
[7 empty] | Part Time Non-Benefit Eligibile | 11.5% |  | #REF! |  | #REF! |  | #REF! |  | #REF!
#REF!

## Non-Helmet
Analysis of Open Source COVID-19 Pandemic Ventilator Projects
2020-03-24 | Public Invention | https://www.pubinv.org |  |  | Home Repo: |  | https://github.com/PubInv/covid19-vent-list
Link to definition of evaluation criteria: | https://docs.google.com/document/d/e/2PACX-1vRl9yZ27KvslftcNvweHgH1A81pO8gHL62TWpY_VY-UELWdK9x-4-3hNw3DbkemClzExPsg8RfnxilP/pub
Project Name | Project Link | Openness | Buildability | Community Support | Functional Testing | Reliability Testing | COVID-19 Suitability | Clinician Friendly | Average | Date Last Evaluated | Point of Contact | Drive | Notes
Medtronic Puritan Bennett (PB) 560 | http://newsroom.medtronic.com/news-releases/news-release-details/medtronic-shares-ventilation-design-specifications-accelerate | 4 | 2.5 | 4 | 5 | 5 | 4 | 5 | 4.21 | 2020-04-01 |  | Pump | Medtronic is a major manufacturer. Release appears to be detailed and complete. Open license is time-limited but good for the pandemic. May be hard to build outside of major manufacturing. Costs of bill of materials and supply chain currently unknown. Thanks to Medtronic for doing this! May not be suitable for ARDS: https://medium.com/@RobertLeeRead/medtronic-open-source-ventilator-does-not-meet-uks-rapidly-manufactured-ventilator-systems-81947c72a7ac
Ambovent | https://1nn0v8ter.rocks/AmboVent-1690-108 | 4.5 | 4 | 4 | 4 | 3 | 4 | 4 | 3.93 | 2020-04-01 | dreliram@gmail.com | AmbuBag | https://tinyurl.com/t2rp2y8 : Made open source on April 2nd. Extremely impressive documentation shown privately. Github org (no repos yet): https://github.com/AmboVent/AmboVent. No license by normal by US legal approach, but clearly dedicated to public domain in verbiage. Thanks to the AmboVent team and Isreal!
Rice OEDK Design: ApolloBVM | http://oedk.rice.edu/apollobvm/ | 5 | 4 | 4 | 2.5 | 2.5 | 3 | 2.5 | 3.36 | 2020-04-03 | amy.k@rice.edu | AmbuBag | Note: Requires registration, but apparently open. Designs are CC-BY, code is GPL in repo: https://github.com/apollobvm/apollobvm 2020 effort informed by older effert, POC is E. D. of OEDK. Open sourced 04/03. Mechanical files are not in repo but downloadable in a zip file once you register.
Low-Cost Open Source | https://github.com/jcl5m1/ventilator | 5 | 4 | 4 | 3 | 0 | 2 | 2 | 2.86 | 2020-03-21 | https://github.com/jcl5m1/ventilator/issues, raise an issue | Pump | Has a good repo and community behind it with active projects
VentilAid | https://www.ventilaid.org/ | 5 | 4 | 4 | 3 | 0 | 2 | 2 | 2.86 | 2020-03-21 | media@urbicum.com | AmbuBag | Great open source project that has amazing documentation that could be built to completion
Open Source Ventilator Project | https://simulation.health.ufl.edu/technology-development/open-source-ventilator-project/ | 3.5 | 2.5 | 4 | 2 | 2 | 3.5 | 2.5 | 2.86 | 2020-03-21 | https://simulation.health.ufl.edu/technology-development/open-source-ventilator-project/contact-us/ |  | Best evidence of reliability. Seems very open, but has no license. Requires close watching!
OpenVentilator (PopSolutions) | https://www.popsolutions.co/en_US/openventilator | 5 | 3.5 | 4 | 2 | 1 | 2 | 2 | 2.79 | 2020-03-30 | contact@openventilator.io | Bellows | Great, well documented project; current test status hard to ascertian, but plans very good. Seems to use very widely available parts.
DIY-Beatmungsgerät [Respirator] | https://devpost.com/software/diy-beatmungsgerat | 5 | 4 | 3 | 1 | 0 | 3 | 3 | 2.71 | 2020-03-21 | https://docs.google.com/forms/d/e/1FAIpQLScmOmF4pyQQUXM5NEUqx-cSPIST0ftLQsrXMesnWElfhT6UqA/viewform |  | In German, may be misreading. May not have PEEP, but mentions good alarm monitoring, so 3.5 on functional tests.
Rice OEDK Design: (2019) ApolloBVM | https://docs.google.com/document/d/1-DRXnVkJOlDCmvTzh-DgWDxeLSrZTiBYyH0ypzv8tNA/edit | 3 | 3 | 3 | 3 | 2 | 3 | 2 | 2.71 | 2020-03-21 | amy.k@rice.edu | AmbuBag | Good documented solution from 2019 created for sneior design in university.
Jeff Ebin's Prototpye | https://www.ebcore.io/?fbclid=IwAR3_SOqqJsohJ4rmR3tQVYS26_RNqx6A1nakRvklNlPNpJ4KS44iOL_a0RU | 5 | 4 | 3 | 1 | 0 | 2 | 3 | 2.57 | 2020-03-21 | jeffrey@ebcore.io |  | Created by MD, Worth a look
A low oxygen consumption pneumatic ventilator for emergency construction during a respiratory failure pandemic | https://onlinelibrary.wiley.com/doi/full/10.1111/j.1365-2044.2009.06207.x | 4 | 3 | 2 | 4 | 1.5 | 1 | 2 | 2.5 | 2020-03-21 | j.dingley@swansea.ac.uk |  | Very good research topic with layed out plans for pressure controls
CITI-OpenLung | https://github.com/emersonmoretto/CITI-OpenLung | 5 | 2 | 4 | 3.5 | 0 | 3 | 0 | 2.5 | 2020-03-21 | https://docs.google.com/forms/d/e/1FAIpQLSfJmSH-XnRPcUmkUvS9ZdUzZxjUVUDR-olO96t6SzpphNBRpQ/viewform |  | Developed a prototype that has undergone testign at a local university. They seem to provide updates daily, and are working on instructions for assembly. In Portuguese.
MUR - Minimal | mur-project.org | 3 | 4 | 3 | 3 | 0 | 2 | 2 | 2.43 | 2020-03-30 | contact@mur-project.org |  | Have a functioning prototype with a videod demonstrating its function.
Protofy Team OxyGEN | https://oxygen.protofy.xyz/ | 4 | 4 | 4 | 1 | 0 | 2 | 1 | 2.29 | 2020-03-21 | https://www.oxygen.protofy.xyz/contact |  | Easy mechanical build
Electric Blower Based Portable Emergency Ventilator | https://digitalcommons.usu.edu/cgi/viewcontent.cgi?referer=https://www.google.com/&httpsredir=1&article=1016&context=spacegrant&fbclid=IwAR1EtJVcxXm82PjGWFCA0t7H_MxNVjuseAePRfxNORr9h4ZQLQ9sNdQjXhc | 3 | 3 | 2 | 3 | 1 | 2 | 1 | 2.14 | 2020-03-21
The Open Ventilator | en.theopenventilator.com | 3 | 2 | 3 | 2 | 0 | 2.5 | 2 | 2.07 | 2020-03-30 | info@theopenventilator.com |  | Provides a video of the functioning prototype on a lung, along with controls.
BlueVent3d | facebook.com/BlueVent3d | 3 | 3 | 4 | 3 | 0 | 0 | 1.5 | 2.07 | 2020-04-02 | facebook.com/BlueVent3d |  | They have several videos showing the functioning prototype, provide design files and a BOM.
MIT 2010 (Husseini et al) | https://web.mit.edu/2.75/projects/DMD_2010_Al_Husseini.pdf | 2 | 2 | 2 | 3 | 1 | 2 | 2 | 2 | 2020-03-28
MIT E-Vent | https://e-vent.mit.edu/ | 2 | 1 | 2.5 | 3 | 1 | 3 | 1 | 1.93 | 2020-03-21 |  | AmuBag | Apparently tested on a live porcine model
Saving Babies' Lives Starts With Aquarium Pumps And Ingenuity | https://www.npr.org/sections/health-shots/2014/01/03/259436844/saving-babies-lives-starts-with-aquarium-pumps-and-ingenuity | 1 | 1 | 2 | 5 | 0 | 0 | 4 | 1.86 | 2020-03-21 |  | Pump | CPAP for infants
RespiraWorks | https://docs.google.com/document/d/1Dz7eMgXowFBtBA_0PKzfAXweHnNMbGlIAXPshCbI2Vk/edit?usp=sharing | 4 | 2 | 3 | 0 | 0 | 4 | 0 | 1.86 | 2020-03-21 | ethan@respira.works |  | Good document for design, not finished, a work in progress
Gtech Ventilator | https://www.gtech.co.uk/ventilators | 4 | 4 | 3 | 0 | 0 | 2 | 0 | 1.86 | 2020-03-30 | ventilator@gtech.co.uk |  | They provide design files and instructions as well as assembly and calibration videos. Not sure if entirely open.
VentilatorPAL | https://freebreathing.org/ | 2 | 1 | 3 | 1.5 | 1 | 1.5 | 2 | 1.71 | 2020-03-21 | https://freebreathing.org/ |  | Pre-orderable. Large but somewhat opaque team, little test information.
Open source ventilator Pakistan | https://docs.google.com/spreadsheets/d/1PfxblJfeVo9WGIUSiz_yQkhYF2OCvRRPal1zllWFMHY/edit#gid=1040995025 | 3 | 1 | 2 | 0 | 0 | 3 | 3 | 1.71 | 2020-03-21 | https://docs.google.com/spreadsheets/d/1PfxblJfeVo9WGIUSiz_yQkhYF2OCvRRPal1zllWFMHY/edit#gid=1746079259 |  | Well documented, large community, but in design phase
CoronavirusMakers | https://gitlab.com/coronavirusmakers | 3 | 2 | 3 | 0 | 0 | 2 | 1 | 1.57 | 2020-03-21 | https://www.coronavirusmakers.org/index.php/es/54-telegram-3 |  | In Spanish, can only read a little, may have more
The Pandemic Ventilator (older) | https://www.instructables.com/id/The-Pandemic-Ventilator/ | 3 | 4 | 1 | 1 | 0 | 1 | 1 | 1.57 | 2020-03-21
Open Source Ventilator - OpenLung BVM Ventilator | https://gitlab.com/open-source-ventilator/OpenLung | 5 | 1 | 4 | 0 | 0 | 0 | 1 | 1.57 | 2020-03-21 | https://opensourceventilator.ie/register
Cuirass-Ventilator | https://github.com/jps2000/Cuirass-Ventilator | 3 | 2 | 3 | 2 | 0 | 0 | 1 | 1.57 | 2020-03-28 | https://github.com/jps2000/Cuirass-Ventilator/issues |  | Cuirass type ventilator, clinical suitability unevaluatable
YACoVV - Yet Another (SARS-)CoV(-2)Ventilator | https://github.com/auenkind/YACoVV | 2 | 1 | 2.5 | 1 | 0 | 3 | 1 | 1.5 | 2020-03-21 |  |  | Fluid pressure based solution
IMPROV: Inexpensive Maker-Made Piston-Respiratory Open-Source Ventilator | https://www.izutadesign.com/ | 4 | 2 | 3 | 1.5 | 0 | 0 | 0 | 1.5 | 2020-03-30 | Design@izutadesign.com |  | Most of the documents they provide describe the design rather than specify it. They performed some functional tests on a piston system rather than a lung or lung simulation.
Ad Hoc Ventilator | https://20100.be/ad-hoc-ventilator/ | 2.5 | 2 | 2 | 1 | 0 | 2 | 1 | 1.5 | 2020-04-03 | https://20100.be/contact-2/ |  | Provides a video of the working device. Essentially an automatic ambu-bag pump
MIT Low Cost Ventilator | https://github.com/RuairiSpain/openVentilator | 3 | 2 | 2 | 1 | 0 | 0 | 2 | 1.43 | 2020-03-21
Dr. Mujeeb ur Rahman design | http://www.technologyreview.pk/pakistani-engineer-braves-tragedy-to-develop-low-cost-ventilator/ | 1 | 1 | 2 | 1 | 0 | 2 | 3 | 1.43 | 2020-03-21
Hackaday Rex Ventilator V1 | https://github.com/hackadayrex/ventilator | 2.5 | 2 | 2.5 | 1 | 0 | 1 | 1 | 1.43 | 2020-03-21 | hackadayrex@gmail.com |  | Great photos and video but no reproducible plans.
RespiraWorks | https://respira.works/ | 1 | 1 | 3 | 3 | 0 | 2 | 0 | 1.43 | 2020-04-02 |  |  | Has closed-loop control, not open source right now. Built own flow measurement! Highly competent leader.
Automatic ambu ventilator | https://github.com/SebaGaggino/rees31 | 3 | 2 | 2 | 0 | 0 | 1 | 1.5 | 1.36 | 2020-03-29
Pandemic Ventilator | https://www.cbc.ca/news/canada/london/pandemic-ventilator-coronvirus-hospitals-1.5493830 | 1 | 1 | 2 | 2 | 0 | 2 | 1 | 1.29 | 2020-03-21 |  | Pump | Inventor says may open source
Open Ventilator Project | https://docs.google.com/document/u/1/d/e/2PACX-1vSrCNygk8LHYIgvpu8tXlVVZrSniwSC7HSBAhMUoOsnFckagZwPCVyNStLS7wBVwR9K0gvYXtRynPKx/pub | 3 | 2 | 4 | 0 | 0 | 0 | 0 | 1.29 | 2020-03-30
OpenVentilator | https://github.com/OpenVentilator/OpenVentilator | 3 | 2 | 3 | 0 | 0 | 1 | 0 | 1.29 | 2020-04-03 |  |  | Not built yet, seems to be design phase
Simple device from www.POMO.cl | https://www.facebook.com/groups/670932227050506/permalink/675264606617268/?app=fbl | 1 | 2 | 3 | 1 | 0 | 0 | 1 | 1.14 | 2020-03-21 | Mserey@pomo.cl
Acute-19 | https://valenciaplaza.com/desarrollan-ventilador-turbina-impresion-3d | 2 | 1 | 2 | 0 | 3 | 0 | 0 | 1.14 | 2020-03-30 | info@acute19.com |  | Reports testing for 48 hours on various lung models. Does not specify parameters or provide details. Will be tested on sheep in the nearby future. Provided website rather empty acute19.com/blog/
COVID19 Respirador (Vaccarini) | https://github.com/vaccarini/covid19respirador | 3 | 2 | 3 | 0 | 0 | 0 | 0 | 1.14 | 2020-03-30 | martin.vaccarini@gmail.com
The Breathing Project | https://www.uni-marburg.de/de/fb13/halbleiterphotonik/the-breathing-project/the-breathing-project-1 | 1.5 | 1 | 3 | 2.5 | 0 | 0 | 0 | 1.14 | 2020-03-30 | breathing.project@physik.uni-marburg.de |  | Provide videos of prototype testing. Controls are visible, but function not demonstrated. State an eventual goal of having the design be public, but have not published any open designs or instructions yet.
Cuirass Ventilator the DIY way | https://www.youtube.com/watch?v=pvrUQCMa3a8&feature=youtu.be | 2 | 2 | 1 | 1 | 0 | 0 | 1 | 1 | 2020-03-21 |  | Cuirass
1M Ventilators | https://1mventilators.com/ | 2 | 1 | 3 | 0 | 0 | 1 | 0 | 1 | 2020-03-21 | https://1mventilators.com/community/ |  | Repo: https://github.com/1M-Ventilators, just beginning
MVP | https://drive.google.com/drive/folders/1XTNc5_pEWI1ATr-WRFbC4TW5NCqKutzC?usp=sharing | 4 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 2020-03-29 | yoram@scope5.com
Open Source Ventilator Ireland | https://opensourceventilator.ie/ | 4 | 2 | 0 | 0 | 0 | 0 | 1 | 1 | 2020-03-28 | opensourceventilatorireland@gmail.com
Breating Aid | https://www.breathing-aid.org/homeen | 2 | 2 | 2 | 0 | 0 | 0 | 1 | 1 | 2020-03-30 | covid@nemos.org |  | Seems to be CPAP for multiple patients; could be useful as relief for equipment shortage for non-ARDS patients?
Low-Cost-Medical-Ventilator | https://github.com/snipe75/low-cost-medical-ventilator | 2.5 | 2 | 1 | 1.5 | 0 | 0 | 0 | 1 | 2020-04-02 |  |  | Don't have build plans but provide their functional code. Works as an ambu bag-type ventilator
Pandemic Ventilator Project | https://panvent.blogspot.com/2008/02/test-of-pandemic-ventilator-with.html | 2 | 1 | 1 | 2 | 0 | 0 | 0 | 0.86 | 2020-03-28
Mechanical Ventilator Milano (MVM) | https://arxiv.org/pdf/2003.10405.pdf | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 0.86 | 2020-03-21
OxVent | https://oxvent.org/ | 1 | 1 | 1 | 3 | 0 | 0 | 0 | 0.86 | 2020-03-28 | https://oxvent.org/contact/ |  | Could be great, but almost no information
Illinois RapidVent | https://rapidvent.grainger.illinois.edu/ | 1 | 1 | 1 | 0 | 0 | 2 | 1 | 0.86 | 2020-03-30 | Recently added, does not appear to be open, non-Ambubag design
Automatic Rescucitator | https://www.thingiverse.com/thing:4239988/files | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0.86 | 2020-03-30 | https://www.frax3d.com/contactar |  | Contact through form on https://www.frax3d.com/contactar. No specifications, mostly just provides 3D models of parts.
Open Source Covid-19 Ventilator Canada | https://open-source-covid-19-ventilator-canada.mn.co/ | 2 | 1 | 3 | 0 | 0 | 0 | 0 | 0.86 | 2020-03-30
Vortran-Type Pneumatic Ventilator | https://docs.google.com/document/d/1R25O2mKT3TfSjXTmheGEevIk6rTJ49o-sGCJU3QqP3M/edit# | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0.86 | 2020-03-30 |  |  | Based off of a commercial design
Low-Cost Automated Emergency Ventilator | https://gitlab.com/PerAsperaAdAstra/lcaev | 1 | 1.5 | 2 | 1 | 0 | 0 | 0 | 0.79 | 2020-03-30 |  |  | Only understandable by view design log: https://gitlab.com/PerAsperaAdAstra/lcaev/-/wikis/Design%20Log
Low-Cost Ventilator Wins Sloan Health Care Prize | https://www.medicaldesignandoutsourcing.com/low-cost-ventilator-wins-sloan-health-care-prize/ | 1 | 1 | 1 | 0 | 0 | 2 | 0 | 0.71 | 2020-03-21 |  |  | Very little information
Open Breath Italy | https://www.openbreath.it/en/ | 1 | 1 | 2 | 0 | 0 | 1 | 0 | 0.71 | 2020-03-21 | info@openbreath.it |  | Not enough information to evaluate well
Projecto EAR Celso | https://github.com/RespiradorHacker/Projeto-EAR-Celso | 2 | 1 | 0 | 0 | 0 | 0 | 2 | 0.71 | 2020-03-21 | https://github.com/RespiradorHacker/Projeto-EAR-Celso
Project Open Air | https://www.projectopenair.org/ | 1 | 1 | 0 | 0 | 0 | 0 | 3 | 0.71 | 2020-03-28 | openair@lift.com.pt
LEITAT1 Respirator | https://www.3dnatives.com/en/3d-printed-respirator-230320205/ | 1 | 2 | 2 | 0 | 0 | 0 | 0 | 0.71 | 2020-03-30 | lluis.lopez@zfbarcelona.es
Respirador RESP19 | https://www.gofundme.com/f/respirador-resp19?utm_medium=chat&utm_source=whatsapp-visit&utm_campaign=p_lico+share-sheet | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 0.71 | 2020-03-30 | https://www.gofundme.com/f/respirador-resp19?utm_medium=chat&utm_source=whatsapp-visit&utm_campaign=p_lico+share-sheet |  | Thus far it's just an idea
OperationAIR | https://www.tudelft.nl/onderwijs/opleidingen/masters/technical-medicine/msc-technical-medicine/operationair/ | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 0.71 | 2020-03-30 | info@operationair.org |  | State that the design and information will be published open source
CoroVent | https://www.corovent.cz/ | 1 | 1 | 3 | 0 | 0 | 0 | 0 | 0.71 | 2020-04-02 | simon.rakosnik@rubikonpr.cz |  | They state that the device is on the prototype phase, and that they are testing components on lung models. They claim to publish the complete documentation, but thus far they don't appear to be open
Ventilador Foscal y Unab | https://www.vanguardia.com/area-metropolitana/bucaramanga/esta-listo-el-primer-ventilador-mecanico-fabricado-100-en-santander-BG2181982 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0.57 | 2020-03-30 |  |  | Very little information, provides a video showing the device pumping
Army-Type Pneumatic Ventilator | https://docs.google.com/document/d/1hRfHGBVZNBeAJTYrPazY6ElMNkyrgIt9aFnz_eUzQG0/edit#heading=h.zd4mdan4zmci | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0.57 | 2020-03-30 |  | AmbuBag | In proposal phase
License: | CC0
Copyright Robert L. Read, Keeshan Patel, and Juan E. Villacres Perez, 2020